## Slide 1: Common Errors and Troubleshooting
Notes: <number>

## Slide 2: Learning Objectives
The LT will know the factors responsible for inaccurate results
The LT will know the common errors
The LT will know how to detect errors resulting in inaccurate results
The LT will know what corrective and preventive actions need to be taken to correct and avoid occurrence of errors
The LT will be able to investigate and trouble shoot the factors responsible for errors
Notes: <number>

## Slide 3: Factors responsible for some common errors
The quality of blood/serum sample
The quality of HIV test kits employed
The environmental conditions (temperature and humidity during storage of kits and performance of test)
The quality status of equipment: maintenance, calibration, etc.
Quality of performance of test (human error)
Notes: <number>

## Slide 4: Other errors
Error in transcription of results (clerical errors)
Non reproducible results, particularly in borderline cases (gray zone cases)
Inter-laboratory variations in test results
Notes: <number>

## Slide 5: Methods to identify common problems
Review protocol/SOP of testing
Check for all parameters which assure quality of HIV kits
Verify whether environmental conditions were met
Check parameters for quality of equipment used
Check for the quality of sample
Examine for any clerical/transcription errors
Examine for possibility of use of reagents from different kits
Examine for possibility of carryover of reagents between adjacent wells
Notes: <number>

## Slide 6: Specimen errors and preventive action
Insufficient volume
Hemolyzed
Pipetting error due to thickening of serum
Bacterial contamination in sample
Aliquoting labeling error
Repeated freeze thawing affects samples, particularly borderline positives
Collect 3ml of blood
Collect a fresh sample and test. If fresh sample is not possible, record sample hemolyzed in result
Record sample as thick in result.
Collect in sterile containier and refrigerate at 2-8º C (one week), -20º C for longer period.
Label carefully. Store properly
Do not freeze thaw repeatedly
Notes: <number>

## Slide 7: Clerical errors and preventive action
Labeling error, mixing of sample ID, wrong history
In result printout the well number and sample number do not match
Error during translation of result from worksheet to report form
Reporting errors
Set up log books/records properly. Collect second sample from same client
Transcribe the results from the printout to work sheet carefully. Repeat test in case of doubt. Please staple result printout to lab worksheet
Do supervisory review and check results
Report to correct person in the correct way after post-test counseling
Notes: <number>

## Slide 8: Test kit dependent errors and preventive action
Use of expired kits
Use of reagents from different kits
Sensitivity & specificity of kit not up to specified standards
Deterioration/contamination of one component
Always check for expiration date of kit and its components before use
Do not mix reagents from different kits
Do not use kits which do not meet the laid down quality standards
Do not use contaminated/expired reagents
Notes: <number>

## Slide 9: Performance errors and preventive action
Dilution error
Scratching off the antigen from the test well with pipette tip
Inconsistent technique
Mixing of samples
SOP not followed
Calculate dilution carefully, accurately
Do not allow pipette tip to touch the well surface
Test each sample including QC sample in the same way
Prepare worksheets meticulously, and number wells/samples with care
Follow every step of SOP
Notes: <number>

## Slide 10: Equipment based errors and preventive action
Use of inappropriate pipettes and pipette tips
Temperature sensitive equipment not working
Equipment not maintained properly and not calibrated as recommended
Use correct volume pipettes and corresponding tips
Maintain daily temperature calibration logs for all such equipment
Maintain and calibrate all equipment as recommended (ELISA Reader, washer, centrifuge, pipettes, etc.)
Notes: <number>

## Slide 11: Environment dependent errors and preventive action
Extremes of temperature and humidity affect enzyme reaction
Drying up of ELISA plates
Control environment. Use A/C. Calibrate incubator & water bath
Maintain humidity. Do not use fan to avoid drying of plates.
Notes: <number>

## Slide 12: Causes of non-reproducible results and preventive action
Mislabeling of sample
Specimen deterioration
Borderline reactors
Carelessness/ignorance LT
Use of obseleteSOP
Collect another sample & repeat test
Take out another aliquot of sample/collect another sample and test in parallel to resolve
Follow up repeat testing to confirm
Vigilance/supervision and training of LT
Use current, valid SOP
Notes: <number>

## Slide 13: Records
Maintain records:
Of each inaccurate result
The cause of any error detected
The corrective actions taken to resolve the error and arrive at the correct result
How the final correct result was obtained
Do not report inaccurate results
The preventive actions planned to prevent repetition of errors
Notes: <number>