## Slide 1
Use and Care of Equipment Available at ICTCs
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 2: Learning Objectives
At the end of this module, you will:
Know your responsibilities related to equipment
Routinely monitor the temperatures of refrigerators and freezers
Confirm that auto pipettes deliver specified volumes accurately
Properly use and maintain centrifuges
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 3: Common equipment used at ICTC
<number>
Centrifuge
Refrigerator
& freezer
Separates cells from serum/ plasma
Pipette
Collects or transfers specimen to test device
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 4: Functioning Equipment is Vital to Quality Service
<number>
Produces
reliable
test results
Lowers
repair costs
Prevents
delays in
testing
Maintains
productivity
Total
Quality
Functioning
Equipment
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 5: Management Responsibilities (Hospital)
Before equipment use:
Update laboratory equipment inventory record
Ensure that manufacturers’ manuals/instructions for all equipment to be available at user site
Establish maintenance program (AMC) including routine function checks (Calibration and temperature logs) and trouble-shooting
Copy of contracts to be available at user sites
Establish maintenance log (Details of breakdowns and repair, time lines action taken)
Develop and implement written protocols for operating procedures
Provide training for all operators
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 6: Responsibilities: Lab Supervisors
Follow written operational procedures
Conduct routine maintenance, including function checks
Take corrective actions
Keep records
<number>
Do not use
malfunctioning
equipment
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 7: Function Checks to Verify that Equipment is Working Properly
Performed routinely as directed
  Daily, weekly, monthly
  After adjustment or repair
Example:
  Monitoring refrigerator temperatures-daily
  Verifying pipette accuracy-quarterly
  Checking centrifuge speed-quarterly
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 8: Refrigerator and Freezer: Use and Care
Keep organized (storing samples and kits)
Periodically clean inside and outside
Ensure door is completely sealed when closing
Ensure that area is well ventilated
Ensure that area is free of rodents
CAUTION! – DO NOT store food items or beverages in laboratory refrigerator or freezer
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 9: Refrigerator & Freezer: Temperature Checks
Monitor daily
Refrigerator: 2oC to 8oC
Freezer: -20oC, -40oC, or -80oC as the case may be
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 10
<number>
Refrigerator & Freezer: Temperature Log
Notes: 2005
Module 11: Use and Care of Equipment
<number>
This form is an example of what should be used to record and monitor refrigerator/freezer temps.
A copy is included in the participant manual.

## Slide 11: Types of Pipettes
Precision pipettes (Not disposable)
  Precise and accurate volumes-Fixed volume or adjustable volume to deliver over a certain range of volume
  Use disposable, single-use, pipette tips
Graduated plastic bulb pipettes (Disposable)
  Dispenses approximate volume
  Easy to use
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 12: Pipette: Use and Care
Select the appropriate pipette for the volume required
Ensure that the pipettor, tips, and specimen are at the same temperature
Firmly attach tip
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 13: Pipette: Use and Care (Cont’d)
Hold the pipette vertically when aspirating
Place tip just below the sample
Avoid air bubbles
Discard contaminated tips in appropriate container after completion of task
<number>
DO NOT RE-USE
  - Pipette tips
  - Graduated plastic bulb pipettes
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 14: Precision Pipettes Require Performance Checks
Performed periodically
Required supplies:
  Pipette
  Pipette tips
  Analytical balance
  Weigh boats
  Distilled or deionzed water
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 15: Pipette: Steps for Checking Reproducibility
Pipette a series of 10 samples into a weigh boat on an analytical scale
Record weight of each sample to calculate calibration results
Verify calculated results are within limits
    Range Max/Min Accuracy
    10 µl + 1.0 µl 10%
    100 µl + 10.0 µ l 10%
    200 µl + 20.0 µ l 10 %
If the results are not within limits, remove from service until appropriate adjustment can be made
Decontaminate pipette and scale after use
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 16: Pipette: Troubleshooting
Tip(s) incorrectly attached | Attach firmly
Foreign articles between the tip and cone O-ring damaged | Clean tip cones Change the O-ring
Incorrect operation Tip incorrectly attached | Follow manufacturer’s instructions carefully Firmly attach tip
Leakage
Problem
Potential Cause
Action
Inaccurate dispensing
Notes: 2005
Module 11: Use and Care of Equipment
<number>
The slide is self explanatory.

## Slide 17: Centrifuge: Use and Care
Always operate with the lid closed
Balance contents before turning on
Check for vibration
Do not open the lid until the rotor has come to a complete stop
Keep lids on tubes when spinning
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 18: Centrifuges: Function Checks
Proper balance
Lubrication
Rotor function
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 19: Centrifuge: Routine Maintenance
Clean interior, condenser coils, fan, and screens
Investigate unusual noises or vibrations
Inspect for evidence of wear, cracks in fitting, corrosion, uneven wear, or signs of fatigue:
  Head, shaft head and coupling
  Rotor
  Brushes and bearings
  Power supply
  Motor and lubricant
  Gaskets, seals, mounts and lubricants
Calibrate speed by accredited centre/purchase tachometer and caliberator
(The service engineer performs all the above mentioned checks)
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 20: Centrifuge Safety
Increase the speed slowly until optimal speed is reached
Disconnect the centrifuge from the electrical source before preventive maintenance, cleaning or inspection
Take caution when removing spills and broken specimen tubes after a run
If tubes are broken, keep the door closed and allow to sit undisturbed for 30 minutes before attempting to clean
Use tweezers to remove broken glass
<number>
Notes: 2005
Module 11: Use and Care of Equipment
<number>

## Slide 21
<number>
Keep a Log for All Maintenance Activities
Notes: 2005
Module 11: Use and Care of Equipment
<number>