## Slide 1
Total Quality Management And Quality Assurance In Laboratories
Notes: <number>

## Slide 2: Learning Objectives
Will know the various components of total quality management .
The Lab Staff will understand the importance of quality assurance.
Will understand mean, median and standard deviation.
Will understand IQC and EQA.
Will understand importance of accuracy of result and how to achieve it.
Notes: <number>

## Slide 3: Quality Management System Objectives
Goal to minimize random and systematic errors
Ensure the reliability, accuracy, reproducibility and precision of observations/results of the test
Reduce cost incurred on account of errors
Notes: <number>

## Slide 4: Context
For a laboratory to be respected as a testing facility, it must be a laboratory that produces excellent (accurate, timely) results, without error within the limitations of the assay system used. If good IQC and quality assurance programme are diligently practiced, the laboratory will be able to meet the needs and expectations of the medical/consumer community by producing highly accurate, reliable and reproducible results.
Total quality management [Includes the quality management, lab management, quality assurance (IQC and EQA)]
Notes: <number>

## Slide 5: Total Quality Management (TQM)
Refers to a comprehensive organizational and lab. approach that is focused on procedures to continually improving the quality and efficiency of processes vide which the laboratory operates.
TQM assists QAP by maximizing efficiency of detection of an error and finding methods to resolve the same.
TQM continually seeks to review and re-evaluate the effectiveness of QC and QA programmes. By internal and external audit.
TQM is a flexible approach which aims to improve QC and QA systems already in place (identify error-modify/eliminate).
TQM is not above QAP but is an adjunct to QAP
TQM monitors the QC/QA programme, technical and administrative considerations which indirectly influence the quality and efficiency of laboratory operations
Notes: <number>

## Slide 6: The Quality System
Notes: <number>

## Slide 7: Elements of TQM
Organizational commitment, provision & procurement
Laboratory management
Documentation & document control
Trouble shooting , preventive and corrective actions
Internal and external audit
Validation of measurements and observations
Feedback to the concerned lab and management
Quality manual &Safety manual including PEP
Transferability/commutability of new technologies
External Quality assessment (EQA)
Notes: <number>

## Slide 8: Philosophy of Quality Assurance Program
What moves needs to be trained
What does not move needs to be calibrated
What happens needs to be documented
What is not documented never happened
Notes: <number>

## Slide 9: Quality Assurance (QA)
Defined as the overall programme/operations in the lab that ensure that the correct result is achieved in a standard reproducible and traceable manner every time a test is performed
QA includes every parameter affecting the results and final reporting i.e. right from collection, transport, logging-in, processing of specimen to interpreting and reporting of result to the correct person in proper manner in a timely manner.
QA for HIV testing also includes factors like:
    Counseling, informed consent, confidentiality
    Inspecting specimens;
    Using the most reliable tests;
    Reviewing transcriptional measures;
    Verifying final reports, etc.
    Reporting result in a timely manner
    Assurance that the results are reported to the appropriate individuals (counseling, confidentiality);
Notes: <number>

## Slide 10: Factors Affecting Quality Assurance
Personnel (Education, training, re-training)
Infrastructure, environment
Equipment, reagents, kits (AMC, calibration, Regular supply)
Availability of relevant documents (SOP, procedure and safety and quality manuals, requisition, reporting formats, consent forms, work books, Levey Jennings charts, etc)
Participation in EQA
Supervision and auditing
Inventory management
Review of results
Preventive, corrective actions and Troubleshooting
Notes: <number>

## Slide 11: Internal Quality Control (QC) Process
IQC refers to those continual measures that must be included during each assay to ensure that a test is working to the highest standard.
IQC indicates that each run produces results that are as accurate as the limits of the test allow but it does not indicate results are correct.
This is so because the accuracy of result depends on the characteristics and limitations of the test as well as operator characteristic.
IQC does not indicate that the results have been reported properly or reported on the correct patient.
Notes: <number>

## Slide 12: Internal quality control
Internal QC Programme: A programme in which your own laboratory runs an external QC sample in an assay and uses the analyzed results (Levey Jennings Chart) to verify validity of the run
Internal controls (kit controls) = Negative and positive kit controls provided by the manufacturer
External control used for IQC every time the test is performed= A sample put on each run of an assay to continually monitor assay performance and validity of ELISA results
QC sample tested at least once a day to validate results of the rapid tests
Notes: <number>

## Slide 13: Control Serum Samples (For Use As IQC Sample)
Source: pooled sera from laboratory, reference centers, commercial sources
Should be filtered through 0.2 nm bacterial filter
Heat inactivated at 560 for 30 minutes
Preparation of ideal set of controls :
* Strong positive
Serial dilution * Strong negative
* Borderline positive
-Most important to interpret “gray zone” reactives)
Done with negative serum rather than NS (normal salina or PBS
Notes: <number>

## Slide 14: Control Serum Samples……..Cont
Storage of controls
Prepared in quantities to last for 6-12 months
Should be sterile preferably without preservatives
Aliquot the sample and store at –200C
1 aliquote thawed and used for 1 week and stored at 2-80 in between.
Notes: <number>

## Slide 15: IQC Programme Method
QC programme method:
External control: The external QC sample is to be included in each run of an assay
Testing: Testing the sample in all assay runs
Determine Acceptance ranges: Ranges for accepting and rejecting runs of an external QC samples (known OD±.2SD Has to be done for commercially and laboratory made external QC samples
Data collection: Entering the data on a daily basis from each run of the external QC sample on LJ chart to detect outliers (for ELISA only)
Analysis: Calculations and graphing of the cumulative data from QC runs
Report: Interpretation of analyzed data and verifying runs based on kit controls and external QC sample performance
Notes: <number>

## Slide 16: IQC Programme Method - Applicable to ELISA only
 | External Quality Control | 
Determine acceptance ranges |  | 
 | Include in all test runs | 
Data collection | Analysis | Report
Notes: <number>

## Slide 17: IQC Programme Method - Applicable to Rapid Tests
 | External Quality Control | 
 |  | Once daily: EQC and IC (kit controls) samples
 | Include in all test runs | 
 | Analysis | Report
Notes: <number>

## Slide 18: IQC … Cont
Essential elements of QC:
Each test run must include one full set of controls
The controls for each test run must yield result within the limits of the manufacture’s/lab criteria for acceptability and validity of the run
Any run failing in above must be repeated
All test kits/reagents must be used before expiration date.
Physical parameters of the test (temp., etc.) must be met.
Notes: <number>

## Slide 19: Variables That Can Affect Quality Of Results
Educational background, certification and training level of the laboratory personnel
The condition of the specimens
The controls used with the test runs
The method of interpretation of the results
The transcription of the results
The reporting of results
Maintenance of confidentiality
The quality of kits reagents
Quality and status of equipment
Environmental conditions
Notes: <number>

## Slide 20: Errors During Performance
Dilution errors
Scratching of the quoted antigen during sample addition
Improper dropper use
Using improper pipette tips
Inconsistent technique
Use of improper equipment
Improper temperature
Mixing components from different kits
Notes: <number>

## Slide 21: Important Issues for Consideration
Selection of source of specimen for external quality control (ex: blood bank patient, etc)
Method of collection of specimen
Transport and storage of specimen, kits and reagents
Test procedures selected
Reference standards (internal/kit controls, EQC sample) used or not
Reagents (Kits) used (Quality approved)
Equipment and instruments used (AMC and calibration)
Documentation (clerical errors : reporting results, overwriting, etc.)
Environmental conditions
Performance of test as per SOP
Notes: <number>

## Slide 22: Criteria for Rejection of Specimen
Inadequate information
Improper labeling
Insufficient volume
Specimen collected in wrong collection tube
CONTROL OF LABORATORY PROCESS
Error can occur at any time from collection of specimen to reporting of result
Error can occur at any step of the laboratory processes
Application of certain rules/methods/checks are necessary to ensure the correctness of procedures. These include general rules/principles and disease specific/process specific rules.
Contamination/haemolysis
Inappropriate transport and storage
Unknown time delay
Notes: <number>

## Slide 23: Laboratory Management
Laboratory Configuration | Stock/inventory management of reagents and controls | Data management
Equipment layout Workflow, staging and direction Space and electrical requirements | Uninterrupted supply Buffer stock First in-First out Prevent stock outs | Performa requisitioning, working, reporting Records Conveying results
Notes: <number>

## Slide 24: Parameters addressed at time of HIV kit selection
Parameters of HIV test kits
      Sensitivity
      Specificity
      Efficiency
      Positive predictive value
      Negative predictive value
Details of commonly available
Rapid test kits
ELISAs
Combinations of assays based on different principles of testing/different antigens if information is available
Notes: <number>

## Slide 25: Corrective action in case of poor performance
Seek inputs (procedure used, reagents used and SOP) to identify the error
Review of QA and IQC procedure
Review of calibration of instruments if required
Review of materials/kits being used
Retraining and reassessment
 If repeated poor performance, than contract must be reconsidered
Notes: <number>

## Slide 26: External Quality Assessment: Method
Inspection: Inspection and accreditation of laboratory by third party (ex: NRL for SRLs and SRLs for ICTCs)
Information obtained during inspection:
Number and qualifications of personnel
Facilities
SOPs
Evidence of QC/Q assurance from records
Participation in proficiency testing
Records/documentation for biosafety procedures practiced
Reporting procedures
Specific recommendation in past implemented or not
Notes: <number>

## Slide 27: Cont...
Proficiency testing:
NRL sends characterized, pedigreed coded panels to the participating laboratory along with procedures, etc. for testing. The results are analyzed by the NRL and quality assessment is made. Both NRL and participating labs benefit from this exercise.
Benefits of proficiency testing:
Difficult or grey areas of participating laboratory revealed can be corrected
Schemes to educate and motivate workers can be planned
Resources to be provided by NACO
NRL can supplement and support participating labs to improve quality standards to the highest level
Notes: <number>

## Slide 28: Evaluation Of Laboratory Staff
Individual meetings assessing:
Quality of work
Quantity of work
Interpersonal relationships
Safety
Punctuality and absenteeism
Professionalism
Continuing education:
        Training and retrainings programs staff members
        Education through periodic in-service training, guest lectures, circulation of pertinent articles, etc. should be ensured to educate staff on advances in technology, etc.
Notes: <number>